## Slide 1: Expected results
 |  |  | Measured Results |  |  | Example Measurements |  | 
Reference |  | Sampling Rate | SNR (dB) | THD (dB) | Reference settling Vin = 4.9V DC | SNR (dB) | THD (dB) | Reference Droop Vin = 4.9V DC (codes)
 | Data Sheet Typical Specification | 1M | 93 | -108 |  | 93 | -108 | 
1 | REF5050 | 1M |  |  |  | 92.2 | -91.4 | 10.9
2 | REF5050 | 500k |  |  |  | 92.1 | -97.1 | 5.35
3 | REF5050 | 100k |  |  |  | 92.0 | -108.2 | 1.4
4 | REF6050 | 1M |  |  |  | 92.6 | -108.1 | 1.15
5 | REF6050 | 500k |  |  |  | 92.7 | -109.3 | 0.8
6 | REF6050 | 100k |  |  |  | 92.3 | -108.1 | 0.8
Notes: The table summarizes the experiments that we will do throughout this experiment. Your measurements will go under the “measured results” section and they should be comparable to the “example measurements” shown at the right. The experiments 1 to 3 are done with the REF5050, and experiments 4 to 6 are done with the REF6050. The experiments will be done at three different sampling rates, and you will see that performance is limited at higher sampling rates for the ADC using REF5050. To test AC performance we will look at SNR and THD, and to test DC performance we will look at reference settling. Print a copy of this table and fill it out as we progress through the experiments. Let’s get started.

## Slide 2: Connect the hardware
Notes: For this experiment we will first use channel 1 on the precision labs hardware. First, set the jumpers as shown here. Connect the PSI to the PLABS board using the SMA cable, and connect the PHI to the PLABS channel 1 connector. Next, connect the USB cables to the computer. Finally, install OPA320_Goodfilter1 in the socket. Be careful to install this amplifier right side up with the label at the bottom of the coupon card. Pause and connect the hardware.

## Slide 3: Start & Setup the Plabs-SAR-EVM Software
3. The green “HW Connected” and teal “PSI Controls” indicate good hardware communication.
2. Select “SAR EVM”
1. Select “Plabs-Toolkit”
Notes: Now let’s start the software by selecting the “Plabs-SAR-EVM” icon from the “Start > All Programs” menu. Once the software is running you should notice the green “HW connected” message at the bottom of the software. Also if the PSI controls are colored teal the hardware communications is working. Pause and get the EVM software started.

## Slide 4: REF5050: 1Msps, THD, SNR
4. Amplitude = 4.9V
Offset = 2.5V
Frequency = 2kHz
Press power on.
1. Select spectral analysis
3. Press Mark Harmonics
2. Sampling Rate 1Msps
Press “Capture”
Record AC performance
 | fs | SNR (dB) | THD (dB)
Spec. | 1M | 93 | -108
REF5050 | 1M | 92.2 | -91.4
Notes: Now let’s measure the AC performance for the system using the low bandwidth REF5050. First, select “spectral analysis” on the “pages” menu. Next, make sure the sampling rate is 1Msps. Next check the “mark Harmonics” option. This will zoom in on the distortion components once we capture the waveform. Now enter the amplitude, offset, and frequency into the PSI Controls. In this case the amplitude is near full scale at 4.9V, the offset is half full scale at 2.5V, and the frequency is 2kHz. Press the power button to enable the PSI output. Note that it is colored teal when the output is enabled and red when it off. Next press capture and record the AC performance. When using REF5050 at 1Msps you see that THD is significantly degraded compared to the data sheet specifications. In this example the THD is -91.4dB whereas the data sheet specification is -108dB. Also, you can see in the FFT that the harmonics are quite large. Pause and measure the AC performance.

## Slide 5: Hardware change
Notes: For this experiment we will make a few minor changes to the hardware. First disconnect the SMA cable from the PSI. Next change the jumper to DC_IN and make sure that the voltage jumper is set to 4.9V. So for this experiment, the AC signal is disconnected and we connect the on-board 4.9V DC source. Pause and change the hardware.

## Slide 6: REF5050: 1Msps, Settling
 | fs | Droop (LSB)
REF5050 | 1M | 10.9
1. Sampling Rate 1Msps
Number of sets to average to 20.
4. Start test
Min Interset delay (ms) to 100ms.
5. Reference droop
Reference Settling Analysis
Notes: Now we can measure reference settling for the REF5050 at 1Msps. First set the sampling rate to 1Msps. Next, set the number of averages to 20. This will take 20 measurements and average them to minimize the noise. Third set the “minimum interset delay” to 100ms. This is just the idle time between each set of 10,000 samples. Fourth, press “start test”. This will take a little time as we are averaging 20 readings. Once this is complete you will see the “reference droop” in LSB. This is the magnitude of the transient error on the reference. In this case the error is 10.9 LSB. You can see most of this happens in the first 500 samples. Pause and measure reference settling.

## Slide 7: Disconnect USB and Close Software
Notes: Now we are going to switch to another channel. To do this we need to disconnect power to both the PSI and PHI by disconnecting the USB cable. Also, close the software should be closed by clicking the “x” in the corner.

## Slide 8: Connect the hardware
Notes: Now we will switch to channel 2 on the precision labs hardware. First, set the jumpers as shown here. Note that we are changing back to an AC input with the jumper settings. Connect the PSI to the PLABS board using the SMA cable, and connect the PHI to the PLABS channel 2 connector. Next, connect the USB cables to the computer. Finally, install OPA320_Goodfilter1 in the socket. Be careful to install this amplifier right side up with the label at the bottom of the coupon card. Pause and connect the hardware.

## Slide 9: Start & Setup the Plabs-SAR-EVM Software
2. The green “HW Connected” and teal “PSI Controls” indicate good hardware communication.
1. Select “Plabs-SAR-EVM” from “start>All Programs”
Notes: Now we can re-start the software by selecting the “Plabs-SAR-EVM” icon from the “Start > All Programs” menu. Once the software is running you should notice the green “HW connected” message at the bottom of the software. Also if the PSI controls are colored teal the hardware communications is working. Pause and get the EVM software started.

## Slide 10: REF6050: 1Msps, 500ksps, 100ksps: SNR, THD
4. Amplitude = 4.9V
Offset = 2.5V
Frequency = 2kHz
1. Select spectral analysis
3. Press Mark Harmonics
2. Sampling Rate 1Msps, 500ksps, 100ksps
Press “Capture”
Record AC performance
 | fs | SNR (dB) | THD (dB)
Spec. | 1M | 93 | -108
REF6050 | 1M | 92.6 | -108.1
REF6050 | 500k | 92.7 | -109.3
REF6050 | 100k | 92.3 | -108.1
Notes: Now we will repeat the SNR and THD measurements for the REF6050. Remember the REF6050 has a wide bandwidth integrated buffer so we expect the system using this device to have good performance up to the full 1Msps sampling rate. To perform this test first select “Spectral Analysis” under “Pages”. Next select the sampling rate. For this test we will use 1Msps, 500ksps, and 100ksps sampling rates. Let’s start with 1Msps. Next check the “Mark Harmonics” check box. This will zoom in on the distortion components in the FFT. Next configure the PSI. Set the amplitude to 4.9V, the offset to 2.5V, and the frequency to 2kHz. Finally, press capture and record the AC performance. Pause and measure SNR and THD at all three sampling rates.
The table to the right compares the SNR and THD specification to the measured performance at the different sampling rates. As expected, the measured performance is very close to the specified performance for all the different sampling rates tested. This is because the wide bandwidth reference in the REF6050 settles quickly allowing for low distortion signal capture. Next we will look at the DC performance with the REF6050.

## Slide 11: Hardware change
Notes: For this experiment we will make a few minor changes to the hardware. First disconnect the SMA cable from the PSI. Next change the jumper to DC_IN and make sure that the voltage jumper is set to 4.9V. So for this experiment, the AC signal is disconnected and we connect the on-board 4.9V DC source. Pause and change the hardware.

## Slide 12: REF6050: 1Msps, 500ksps, and 100ksps settling
Sampling Rate 1Msps, 500ksps, and 100ksps
3. Number of sets to average to 20.
5. Start test
4. Min Interset delay (ms) to 100
6. Reference droop
 | fs | Droop (LSB)
REF6050 | 1M | 1.15
REF6050 | 500k | 0.8
REF6050 | 100k | 0.8
Select “reference settling”
Notes: Here we will check the reference settling for the REF6050 by applying a DC 4.9V input and measuring the conversions. When we performed the same experiment using the REF5050 the performance was degraded at higher sampling rates. For this experiment first select “Reference Settling” under “Pages”. Next, select the sampling rate. Here we will run the experiment for 1Msps, 500ksps, and 100ksps. Let’s start with 1Msps. Next enter the number of data sets that we will average to 20. Next set the “Interset” delay to 100ms. This is the delay between each set of points being measured. Finally, press “Start test” and measure the reference droop. The figure shown shows the results for 1Msps. Repeat the experiment for 500ksps, and 100ksps. Pause and measure droop at all three sampling rates.
The table to the right shows the settling performance for the REF5050 system. Ideally the droop should be less than 1 LSB. As expected, the measured performance is very close to this target droop error. This is because the wide bandwidth reference in the REF6050 responds quickly to the transient current requirements and holds the reference voltage constant. Now let’s compare all the measured results.